## Summary
Parameter nE, number of bins for the energy space used to calculated the quenched energy
Target relatif standard uncertainty <= |  |  | 0.0005
Radionuclide | Ebeta max /keV | nE | a | N
H-3 | 18.591 | 7000 | 0.303 | 380592.7599
Pu-241 | 20.8 | 7000 | 0.337 | 340173.0769
Ni-63 | 66.98 | 2000 | 0.315 | 105637.5037
C-14 | 156.476 | 1000 | 0.186 | 45218.4361
S-35 | 167.33 | 1000 | 0.208 | 42285.3045
Pm-147 | 224.1 | 1000 | 0.163 | 31573.4047
Ca-45 | 258 | 500 | 0.182 | 27424.8062
Tc-99 | 293.8 | 500 | 0.136 | 24083.0497
Sr-90 | 545.9 | 200 | 0.106 | 12961.3482
Sr-89 | 1495.1 | 50 | 0.0467 | 4732.5263
#DIV/0!
Convergence function f = a/sqrt(N) |  | f, relative standard uncertainty
f = a / sqrt(N)
a = 1.33 / sqrt(Q)
f = 1.33 / sqrt(QN)
Factor | 1.33 | 1.7689
Comparison of TDCRPy (analytical calculation) with the LNHB code TDCR17
Analytical model, L=1 keV-1, kB=1e-5 cm/keV |  |  |  |  | Radionuclide | modèle analytique | modèle Monte-Carlo | u | dev | ref
LS Cocktail | Ultima Gold |  |  |  | H-3 | 0.59436 | 0.59433 | 0.001101 | -2.4611e-05 | 0
lower integration bound | 64.6 eV |  |  |  | Pu-241 | 0.52635 | 0.52688 | 0.0011686 | 0.0005296 | 0
[5 empty] | Ni-63 | 0.83224 | 0.83268 | 0.00095023 | 0.00043915 | 0
Deviation = TDCRPy/TDCR17-1 |  |  |  |  | C-14 | 0.95409 | 95.44% | 0.05431% | 0.00026381 | 0
Radionuclide | Ebeta max /keV | eff_D | eff_T | eff_T/eff_D | S-35 | 0.93955 | 0.93866 | 0.00063107 | -0.00088736 | 0
H-3 | 18.591 | -0.34% | -0.24% | 0.11% | Pm-147 | 0.94773 | 0.94804 | 0.00058916 | 0.00030427 | 0
Pu-241 | 20.8 | 0.14% | 0.8% | 0.66% | Ca-45 | 0.96207 | 0.96061 | 0.00051933 | -0.0014588 | 0
Ni-63 | 66.98 | 0.08% | 0.08% | 0% | Tc-99 | 0.97443 | 97.35% | 0.04249% | -0.00090639 | 0
C-14 | 156.476 | -0.33% | -0.33% | 0.01% | Sr-90 | 0.98644 | 0.98602 | 0.00031323 | -0.00042195 | 0
S-35 | 167.33 | 0.31% | 0.5% | -0.19% | Sr-89 | 0.9956 | 0.99386 | 0.00021944 | -0.0017437 | 0
Pm-147 | 224.1 | -0.6% | -0.87% | 0.27%
Ca-45 | 258 | 0.48% | 0.48% | 0%
Tc-99 | 293.8 | 0.21% | 0.68% | 0.47%
Sr-90 | 545.9 | -0.06% | 0.16% | 0.21%
Sr-89 | 1495.1 | -0.26% | -0.3% | -0.04%

## H-3
The Stochastic implementation of TDCRPy is tested with its analystical implementation in the cases of pure beta emitting radionuclides
Convergence of the results with the number of points considered in the quenching function (parameter nE)
Analytical model, L=1 keV-1, kB=1e-5 cm/keV | [5 empty] | H-3 | Q | 18.591 | keV
nE | eff_D | Err
3 | 77.67% | 30.69%
4 | 73.56% | 23.77%
5 | 70.93% | 19.35%
6 | 69.1% | 16.27%
7 | 67.76% | 14.01%
10 | 65.26% | 9.802%
20 | 62.21% | 4.684%
50 | 60.4% | 1.628%
80 | 60% | 0.9511%
100 | 59.87% | 0.7485%
200 | 59.65% | 0.3726%
500 | 59.52% | 0.1473%
1000 | 59.47% | 0.07282%
2000 | 59.45% | 0.03573%
5000 | 59.44% | 0.01372%
7000 | 59.44% | 0.009403%
10000 | 59.43% | 0.006359%
20000 | 59.43% | 0.002637%
50000 | 59.43% | 0.0004299%
70000 | 59.43% | 0%
reference | 59.43%
recommanded value |  | 7000
Convergence of the result of the stochastic model with the results o the analytical model (N = number of Monte-Carlo trials) | [9 empty] | Analytical model results
Analytical model, L=1 keV-1, kB=1e-5 cm/keV | [9 empty] | eff | TDCRPy | TDCR17 restults | deviation
N | eff_S | u(eff_S) | eff_D | u(eff_D) | eff_T | u(eff_T) | dev S | dev D | dev T | eff_S | 0.57608
10 | 65.61% | 7.947% | 71.67% | 9.377% | 38.21% | 6.875% | 13.89% | 20.59% | 17.23% | eff_D | 0.59436 | 0.5964 | -0.3427%
20 | 64.69% | 5.176% | 68.19% | 6.356% | 36.72% | 6.176% | 12.3% | 14.73% | 12.67% | eff_T | 0.32592 | 0.3267 | -0.2372%
50 | 55.98% | 3.961% | 58.68% | 4.866% | 29.44% | 3.512% | -2.819% | -1.266% | -9.671% | TDCR | 0.54837 | 0.54779 | 0.1059% | 0.65611 | 0.07947 | 0.71675 | 0.093766 | 0.38207 | 0.068746
100 | 54.1% | 2.836% | 54.88% | 3.454% | 28.52% | 2.901% | -6.096% | -7.668% | -12.48% |  |  |  |  | 0.64693 | 0.051761 | 0.6819 | 0.063556 | 0.36721 | 0.061758
200 | 55.25% | 2.121% | 56.61% | 2.523% | 30.98% | 2.148% | -4.095% | -4.76% | -4.955% |  |  |  |  | 0.55984 | 0.039614 | 0.58683 | 0.048665 | 0.29441 | 0.035116
500 | 59.14% | 1.298% | 61.26% | 1.56% | 34.46% | 1.355% | 2.655% | 3.066% | 5.741% |  |  |  |  | 0.54096 | 0.028364 | 0.54878 | 0.034543 | 0.28525 | 0.029008
1000 | 57.06% | 0.9087% | 58.71% | 1.095% | 31.87% | 0.9417% | -0.9524% | -1.223% | -2.219% |  |  |  |  | 0.55249 | 0.021208 | 0.56606 | 0.025225 | 0.30978 | 0.021475
2000 | 57.51% | 0.6413% | 59.28% | 0.7761% | 32.31% | 0.6628% | -0.1724% | -0.2617% | -0.8708% |  |  |  |  | 0.59138 | 0.012979 | 0.61258 | 0.015603 | 0.34464 | 0.013552
5000 | 57.67% | 0.4051% | 59.39% | 0.4892% | 32.49% | 0.4231% | 0.09904% | -0.08364% | -0.3051% |  |  |  |  | 0.57059 | 0.0090873 | 0.58709 | 0.010951 | 0.31869 | 0.0094174
10000 | 57.37% | 0.2888% | 59.1% | 0.3487% | 32.32% | 0.2977% | -0.4194% | -0.5618% | -0.8376% |  |  |  |  | 0.57509 | 0.0064135 | 0.5928 | 0.0077609 | 0.32309 | 0.0066278
20000 | 57.46% | 0.205% | 59.23% | 0.2471% | 32.51% | 0.2112% | -0.2539% | -0.3437% | -0.2416% |  |  |  |  | 0.57665 | 0.0040508 | 0.59386 | 0.0048923 | 0.32493 | 0.0042314
50000 | 57.74% | 0.1292% | 59.55% | 0.1557% | 32.75% | 0.1339% | 0.2277% | 0.1988% | 0.4766% |  |  |  |  | 0.57367 | 0.0028875 | 0.59102 | 0.003487 | 0.32319 | 0.0029775
100000 | 57.61% | 0.09142% | 59.43% | 0.1101% | 32.6% | 0.09435% | -0.004261% | -0.004141% | 0.008895% |  |  |  |  | 0.57462 | 0.0020495 | 0.59231 | 0.0024715 | 0.32514 | 0.0021124
[7 empty] | -100% | -100% | -100% |  |  |  |  | 0.57739 | 0.0012919 | 0.59554 | 0.0015565 | 0.32748 | 0.0013391
[7 empty] | -100% | -100% | -100% |  |  |  |  | 0.57606 | 0.00091418 | 0.59433 | 0.001101 | 0.32595 | 0.00094351
Relative uncertainty due to MC calculation
N | 100000
u_r | 0.1165%
recommanded value |  | 10000
the calculation uncertainty provided by the code in consistent with the observation

## Pu-241
The Stochastic implementation of TDCRPy is tested with its analystical implementation in the cases of pure beta emitting radionuclides
Convergence of the results with the number of points considered in the quenching function (parameter nE)
Analytical model, L=1 keV-1, kB=1e-5 cm/keV | [5 empty] | Pu-241 | Q | 20.8 | keV
nE | eff_D | Err
3 |  | -100%
4 |  | -100%
5 |  | -100%
6 |  | -100%
7 |  | -100%
10 |  | -100%
20 |  | -100%
50 |  | -100%
80 |  | -100%
100 | 53.04% | 0.7844%
200 | 52.84% | 0.3902%
500 | 52.71% | 0.1539%
1000 | 52.67% | 0.07585%
2000 | 52.65% | 0.03699%
5000 | 52.64% | 0.01392%
7000 | 52.63% | 0.009397%
10000 | 52.63% | 0.006211%
20000 | 52.63% | 0.002312%
50000 | 52.63% | 0%
70000 |  | -100%
reference | 52.63%
recommanded value |  | 7000
Convergence of the result of the stochastic model with the results o the analytical model (N = number of Monte-Carlo trials) | [9 empty] | Analytical model results
Analytical model, L=1 keV-1, kB=1e-5 cm/keV | [9 empty] | eff | TDCRPy | TDCR17 restults | deviation
N | eff_S | u(eff_S) | eff_D | u(eff_D) | eff_T | u(eff_T) | dev S | dev D | dev T | eff_S | 0.51917 |  | #DIV/0!
10 | 47.42% | 8.864% | 47.13% | 11.14% | 21.94% | 7.682% | -8.669% | -10.45% | -23.54% | eff_D | 0.52635 | 0.5256 | 0.1424% | 0.47416 | 0.088638 | 0.47132 | 0.11141 | 0.21943 | 0.076817
20 | 64.04% | 4.781% | 67.93% | 5.952% | 34.41% | 5.68% | 23.35% | 29.06% | 19.91% | eff_T | 0.28699 | 0.2847 | 0.8028% | 0.64042 | 0.047813 | 0.67929 | 0.059518 | 0.34414 | 0.056803
50 | 54.3% | 3.85% | 55.23% | 4.729% | 27.73% | 3.935% | 4.599% | 4.94% | -3.366% | TDCR | 0.54524 | 0.54167 | 0.6595% | 0.54305 | 0.038497 | 0.55235 | 0.047286 | 0.27733 | 0.039346
100 | 53.24% | 2.919% | 54.09% | 3.496% | 28.25% | 2.927% | 2.546% | 2.774% | -1.578% |  |  |  |  | 0.53239 | 0.029186 | 0.54095 | 0.034956 | 0.28246 | 0.029271
200 | 49.36% | 2.18% | 49.45% | 2.616% | 26.09% | 2.075% | -4.919% | -6.057% | -9.098% |  |  |  |  | 0.49363 | 0.021802 | 0.49447 | 0.026159 | 0.26088 | 0.020746
500 | 53.68% | 1.394% | 54.65% | 1.652% | 30.49% | 1.393% | 3.398% | 3.821% | 6.229% |  |  |  |  | 0.53681 | 0.013945 | 0.54646 | 0.016519 | 0.30486 | 0.013925
1000 | 52.21% | 0.9805% | 52.93% | 1.159% | 28.84% | 0.9677% | 0.5586% | 0.5635% | 0.4857% |  |  |  |  | 0.52207 | 0.0098048 | 0.52931 | 0.011591 | 0.28838 | 0.0096773
2000 | 52% | 0.6999% | 52.72% | 0.8308% | 28.9% | 0.6824% | 0.1678% | 0.1583% | 0.7109% |  |  |  |  | 0.52004 | 0.0069988 | 0.52718 | 0.0083076 | 0.28903 | 0.0068244
5000 | 51.9% | 0.4432% | 52.62% | 0.5248% | 28.83% | 0.4328% | -0.02782% | -0.02902% | 0.4627% |  |  |  |  | 0.51903 | 0.0044321 | 0.5262 | 0.0052484 | 0.28831 | 0.0043279
10000 | 51.95% | 0.3124% | 52.74% | 0.3703% | 28.74% | 0.3034% | 0.05475% | 0.2063% | 0.1611% |  |  |  |  | 0.51945 | 0.0031243 | 0.52743 | 0.0037025 | 0.28745 | 0.0030342
20000 | 51.48% | 0.221% | 52.14% | 0.2619% | 28.34% | 0.2146% | -0.8377% | -0.9322% | -1.244% |  |  |  |  | 0.51482 | 0.0022104 | 0.52144 | 0.0026188 | 0.28342 | 0.002146
50000 | 51.91% | 0.1397% | 52.62% | 0.1656% | 28.75% | 0.1364% | -0.01281% | -0.03039% | 0.1786% |  |  |  |  | 0.5191 | 0.0013971 | 0.52619 | 0.0016559 | 0.2875 | 0.0013644
100000 | 51.96% | 0.09862% | 52.69% | 0.1169% | 28.75% | 0.09639% | 0.08841% | 0.1006% | 0.1672% |  |  |  |  | 0.51963 | 0.00098618 | 0.52688 | 0.0011686 | 0.28747 | 0.00096387
[7 empty] | -100% | -100% | -100%
[7 empty] | -100% | -100% | -100%
Relative uncertainty due to MC calculation
N | 100000
u_r | 0.09771%
recommanded value |  | 10000
the calculation uncertainty provided by the code in consistent with the observation

## C-14
The Stochastic implementation of TDCRPy is tested with its analystical implementation in the cases of pure beta emitting radionuclides
Convergence of the results with the number of points considered in the quenching function (parameter nE)
Analytical model, L=1 keV-1, kB=1e-5 cm/keV | [5 empty] | C-14 | Q | 156.476 | keV
nE | eff_D | Err
3 | 97.44% | 2.134%
4 | 97.01% | 1.683%
5 | 96.73% | 1.387%
6 | 96.53% | 1.178%
7 | 96.38% | 1.022%
10 | 96.1% | 0.7266%
20 | 95.74% | 0.356%
50 | 95.53% | 0.1265%
80 | 95.48% | 0.07375%
100 | 95.46% | 0.05756%
200 | 95.43% | 0.02799%
500 | 95.41% | 0.01039%
1000 | 95.41% | 0.004784%
2000 | 95.41% | 0.002078%
5000 | 95.41% | 0.0005143%
7000 | 95.41% | 0.0002334%
10000 | 95.4% | 0%
20000 |  | -100%
50000 |  | -100%
70000 |  | -100%
reference | 95.4%
recommanded value |  | 1000
Convergence of the result of the stochastic model with the results o the analytical model (N = number of Monte-Carlo trials) | [9 empty] | Analytical model results
Analytical model, L=1 keV-1, kB=1e-5 cm/keV | [9 empty] | eff | TDCRPy | TDCR17 restults | deviation
N | eff_S | u(eff_S) | eff_D | u(eff_D) | eff_T | u(eff_T) | dev S | dev D | dev T | eff_S | 0.94789 |  | #DIV/0!
10 | 89.89% | 9.476% | 90% | 9.487% | 89.68% | 9.457% | -5.165% | -5.672% | -1.676% | eff_D | 0.95409 | 0.9573 | -0.3349%
20 | 90.75% | 4.555% | 91.75% | 4.882% | 83.88% | 6.713% | -4.261% | -3.836% | -8.036% | eff_T | 0.91212 | 0.9151 | -0.3254%
50 | 88.82% | 3.487% | 89.39% | 3.719% | 82.77% | 4.618% | -6.294% | -6.312% | -9.256% | TDCR | 0.95601 | 0.95592 | 0.009528% | 0.89894 | 0.09476 | 0.89997 | 0.094866 | 0.89683 | 0.09457
100 | 94.95% | 1.369% | 96.02% | 1.374% | 90.02% | 2.325% | 0.1645% | 0.639% | -1.304% |  |  |  |  | 0.9075 | 0.045547 | 0.91749 | 0.048818 | 0.83883 | 0.067134
200 | 94.47% | 1.301% | 94.93% | 1.362% | 91.49% | 1.685% | -0.3359% | -0.5073% | 0.3007% |  |  |  |  | 0.88823 | 0.034868 | 0.89387 | 0.03719 | 0.82769 | 0.046182
500 | 95.53% | 0.7347% | 96.21% | 0.7477% | 92.83% | 0.9368% | 0.7789% | 0.8344% | 1.772% |  |  |  |  | 0.94945 | 0.013694 | 0.96019 | 0.013744 | 0.90023 | 0.023246
1000 | 95.51% | 0.4926% | 96.02% | 0.5154% | 92.46% | 0.6913% | 0.7593% | 0.6448% | 1.365% |  |  |  |  | 0.94471 | 0.01301 | 0.94925 | 0.013619 | 0.91487 | 0.01685
2000 | 94.75% | 0.3818% | 95.22% | 0.4017% | 91.43% | 0.5288% | -0.04157% | -0.1953% | 0.2332% |  |  |  |  | 0.95528 | 0.0073468 | 0.96206 | 0.0074771 | 0.92829 | 0.0093675
5000 | 94.77% | 0.2357% | 95.39% | 0.2454% | 91.2% | 0.3314% | -0.01839% | -0.016% | -0.0178% |  |  |  |  | 0.95509 | 0.004926 | 0.96025 | 0.0051545 | 0.92457 | 0.0069134
10000 | 94.68% | 0.1664% | 95.33% | 0.173% | 90.97% | 0.2367% | -0.1109% | -0.08651% | -0.2701% |  |  |  |  | 0.9475 | 0.003818 | 0.95223 | 0.0040171 | 0.91425 | 0.0052883
20000 | 94.68% | 0.1179% | 95.32% | 0.1225% | 90.97% | 0.1674% | -0.1175% | -0.08977% | -0.2673% |  |  |  |  | 0.94772 | 0.0023572 | 0.95394 | 0.0024536 | 0.91196 | 0.0033135
50000 | 94.76% | 0.07405% | 95.41% | 0.07679% | 91.11% | 0.105% | -0.02636% | 0.002609% | -0.1109% |  |  |  |  | 0.94684 | 0.0016639 | 0.95327 | 0.0017301 | 0.90966 | 0.0023674
100000 | 94.81% | 0.05228% | 95.44% | 0.05431% | 91.21% | 0.07398% | 0.01942% | 0.02765% | -0.003611% |  |  |  |  | 0.94678 | 0.0011792 | 0.95324 | 0.0012248 | 0.90968 | 0.0016739
[7 empty] | -100% | -100% | -100% |  |  |  |  | 0.94764 | 0.00074051 | 0.95412 | 0.00076787 | 0.91111 | 0.00105
[7 empty] | -100% | -100% | -100% |  |  |  |  | 0.94808 | 0.00052276 | 0.95436 | 0.0005431 | 0.91209 | 0.00073984
Relative uncertainty due to MC calculation
N | 100000
u_r | 0.08054%
recommanded value |  | 10000
the calculation uncertainty provided by the code in consistent with the observation

## S-35
The Stochastic implementation of TDCRPy is tested with its analystical implementation in the cases of pure beta emitting radionuclides
Convergence of the results with the number of points considered in the quenching function (parameter nE)
Analytical model, L=1 keV-1, kB=1e-5 cm/keV | [5 empty] | S-35 | Q | 167.33 | keV
nE | eff_D | Err
3 |  | -100%
4 |  | -100%
5 |  | -100%
6 |  | -100%
7 |  | -100%
10 | 94.81% | 0.9195%
20 | 94.37% | 0.4477%
50 | 94.1% | 0.1584%
80 | 94.04% | 0.09241%
100 | 94.02% | 0.0722%
200 | 93.98% | 0.03516%
500 | 93.96% | 0.01303%
1000 | 93.95% | 0.005992%
2000 | 93.95% | 0.002601%
5000 | 93.95% | 0.0006437%
7000 | 93.95% | 0.000293%
10000 | 93.95% | 0%
20000 |  | -100%
50000 |  | -100%
70000 |  | -100%
reference | 93.95%
recommanded value |  | 1000
Convergence of the result of the stochastic model with the results o the analytical model (N = number of Monte-Carlo trials) | [9 empty] | Analytical model results
Analytical model, L=1 keV-1, kB=1e-5 cm/keV | [9 empty] | eff | TDCRPy | TDCR17 restults | deviation
N | eff_S | u(eff_S) | eff_D | u(eff_D) | eff_T | u(eff_T) | dev S | dev D | dev T | eff_S | 0.93292
10 | 0.99689 | 0.0024347 | 0.9998 | 0.0001847 | 0.99089 | 0.0071204 | 6.857% | 6.413% | 11.37% | eff_D | 0.93955 | 0.9425 | 0.3143%
20 | 0.91772 | 0.054528 | 0.90909 | 0.06097 | 0.89798 | 0.066472 | -1.629% | -3.242% | 0.9266% | eff_T | 0.88973 | 0.8942 | 0.5021%
50 | 0.9675 | 0.020506 | 0.97462 | 0.019986 | 0.9483 | 0.024319 | 3.706% | 3.733% | 6.583% | TDCR | 0.94698 | 0.94875 | -0.1869%
100 | 0.91462 | 0.02093 | 0.92119 | 0.022702 | 0.85991 | 0.029035 | -1.962% | -1.954% | -3.352% |  |  |  |  | 0.99689 | 0.0024347 | 0.9998 | 0.0001847 | 0.99089 | 0.0071204
200 | 0.94265 | 0.012892 | 0.94713 | 0.013528 | 0.90918 | 0.017337 | 1.043% | 0.8068% | 2.186% |  |  |  |  | 0.91772 | 0.054528 | 0.90909 | 0.06097 | 0.89798 | 0.066472
500 | 0.93239 | 0.0083947 | 0.94016 | 0.0087582 | 0.88681 | 0.011716 | -0.0568% | 0.06474% | -0.329% |  |  |  |  | 0.9675 | 0.020506 | 0.97462 | 0.019986 | 0.9483 | 0.024319
1000 | 0.93754 | 0.0057177 | 0.94497 | 0.0059223 | 0.89503 | 0.0080203 | 0.4953% | 0.5768% | 0.5957% |  |  |  |  | 0.91462 | 0.02093 | 0.92119 | 0.022702 | 0.85991 | 0.029035
2000 | 0.93123 | 0.0043506 | 0.93741 | 0.0045533 | 0.88886 | 0.005933 | -0.1811% | -0.2274% | -0.0976% |  |  |  |  | 0.94265 | 0.012892 | 0.94713 | 0.013528 | 0.90918 | 0.017337
5000 | 0.9304 | 0.0027565 | 0.93655 | 0.0028917 | 0.88718 | 0.003781 | -0.2703% | -0.3194% | -0.2872% |  |  |  |  | 0.93239 | 0.0083947 | 0.94016 | 0.0087582 | 0.88681 | 0.011716
10000 | 0.92924 | 0.0019569 | 0.93578 | 0.0020464 | 0.88479 | 0.0026919 | -0.3942% | -0.4011% | -0.5555% |  |  |  |  | 0.93754 | 0.0057177 | 0.94497 | 0.0059223 | 0.89503 | 0.0080203
20000 | 0.93058 | 0.0013751 | 0.93719 | 0.0014368 | 0.88711 | 0.0018824 | -0.2506% | -0.2505% | -0.2948% |  |  |  |  | 0.93123 | 0.0043506 | 0.93741 | 0.0045533 | 0.88886 | 0.005933
50000 | 0.93317 | 0.0008445 | 0.93987 | 0.00088244 | 0.88976 | 0.0011733 | 0.02657% | 0.03462% | 0.002881% |  |  |  |  | 0.9304 | 0.0027565 | 0.93655 | 0.0028917 | 0.88718 | 0.003781
100000 | 0.93218 | 0.00060401 | 0.93866 | 0.00063107 | 0.88883 | 0.0008363 | -0.0796% | -0.09445% | -0.1017% |  |  |  |  | 0.92924 | 0.0019569 | 0.93578 | 0.0020464 | 0.88479 | 0.0026919
[7 empty] | -100% | -100% | -100% |  |  |  |  | 0.93058 | 0.0013751 | 0.93719 | 0.0014368 | 0.88711 | 0.0018824
[7 empty] | -100% | -100% | -100% |  |  |  |  | 0.93317 | 0.0008445 | 0.93987 | 0.00088244 | 0.88976 | 0.0011733
[14 empty] | 0.93218 | 0.00060401 | 0.93866 | 0.00063107 | 0.88883 | 0.0008363
Relative uncertainty due to MC calculation
N | 100000
u_r | 0.1044%
recommanded value |  | 10000
the calculation uncertainty provided by the code in consistent with the observation

## Pm-147
The Stochastic implementation of TDCRPy is tested with its analystical implementation in the cases of pure beta emitting radionuclides
Convergence of the results with the number of points considered in the quenching function (parameter nE)
Analytical model, L=1 keV-1, kB=1e-5 cm/keV | [5 empty] | Pm-147 | Q | 224.1 | keV
nE | eff_D | Err
3 |  | -100%
4 |  | -100%
5 |  | -100%
6 |  | -100%
7 |  | -100%
10 | 95.49% | 0.7595%
20 | 95.12% | 0.3701%
50 | 94.89% | 0.1305%
80 | 94.84% | 0.07565%
100 | 94.82% | 0.05901%
200 | 94.8% | 0.02841%
500 | 94.78% | 0.01059%
1000 | 94.77% | 0.004639%
2000 | 94.77% | 0.001721%
5000 | 94.77% | 0%
7000 |  | -100%
10000 |  | -100%
20000 |  | -100%
50000 |  | -100%
70000 |  | -100%
reference | 94.77%
recommanded value |  | 1000
Convergence of the result of the stochastic model with the results o the analytical model (N = number of Monte-Carlo trials) | [9 empty] | Analytical model results
Analytical model, L=1 keV-1, kB=1e-5 cm/keV | [9 empty] | eff | TDCRPy | TDCR17 restults | deviation
N | eff_S | u(eff_S) | eff_D | u(eff_D) | eff_T | u(eff_T) | dev S | dev D | dev T | eff_S | 0.94214 |  | #DIV/0!
10 | 0.86449 | 0.093801 | 0.88324 | 0.093709 | 0.81139 | 0.10191 | -8.242% | -6.805% | -10.45% | eff_D | 0.94773 | 0.9535 | -0.6047%
20 | 0.98151 | 0.013663 | 0.98995 | 0.0091897 | 0.95566 | 0.030964 | 4.178% | 4.454% | 5.475% | eff_T | 0.90605 | 0.914 | -0.8695% | 0.86449 | 0.093801 | 0.88324 | 0.093709 | 0.81139 | 0.10191
50 | 0.95118 | 0.021922 | 0.95147 | 0.023368 | 0.91742 | 0.035059 | 0.9588% | 0.3938% | 1.255% | TDCR | 0.95602 | 0.95857 | -0.2664% | 0.98151 | 0.013663 | 0.98995 | 0.0091897 | 0.95566 | 0.030964
100 | 0.94745 | 0.016997 | 0.95329 | 0.017247 | 0.91317 | 0.023764 | 0.5629% | 0.5865% | 0.7858% |  |  |  |  | 0.95118 | 0.021922 | 0.95147 | 0.023368 | 0.91742 | 0.035059
200 | 0.93508 | 0.013838 | 0.93936 | 0.014456 | 0.89932 | 0.018403 | -0.7502% | -0.884% | -0.7431% |  |  |  |  | 0.94745 | 0.016997 | 0.95329 | 0.017247 | 0.91317 | 0.023764
500 | 0.95563 | 0.0064822 | 0.96277 | 0.0066951 | 0.91998 | 0.0097155 | 1.431% | 1.586% | 1.537% |  |  |  |  | 0.93508 | 0.013838 | 0.93936 | 0.014456 | 0.89932 | 0.018403
1000 | 0.93911 | 0.0059974 | 0.94468 | 0.0062516 | 0.90449 | 0.0078618 | -0.3225% | -0.3225% | -0.172% |  |  |  |  | 0.95563 | 0.0064822 | 0.96277 | 0.0066951 | 0.91998 | 0.0097155
2000 | 0.93984 | 0.0040922 | 0.94587 | 0.0042266 | 0.90224 | 0.0056102 | -0.245% | -0.1971% | -0.4203% |  |  |  |  | 0.93911 | 0.0059974 | 0.94468 | 0.0062516 | 0.90449 | 0.0078618
5000 | 0.94417 | 0.002525 | 0.94941 | 0.0026284 | 0.91029 | 0.0034228 | 0.2147% | 0.1772% | 0.4678% |  |  |  |  | 0.93984 | 0.0040922 | 0.94587 | 0.0042266 | 0.90224 | 0.0056102
10000 | 0.94172 | 0.0017916 | 0.94767 | 0.001859 | 0.90457 | 0.0024709 | -0.04489% | -0.006425% | -0.1638% |  |  |  |  | 0.94417 | 0.002525 | 0.94941 | 0.0026284 | 0.91029 | 0.0034228
20000 | 0.93894 | 0.0013046 | 0.94449 | 0.0013581 | 0.90122 | 0.0017906 | -0.3404% | -0.3418% | -0.5336% |  |  |  |  | 0.94172 | 0.0017916 | 0.94767 | 0.001859 | 0.90457 | 0.0024709
50000 | 0.94178 | 0.00080495 | 0.94735 | 0.00083811 | 0.90536 | 0.001106 | -0.0382% | -0.04089% | -0.07644% |  |  |  |  | 0.93894 | 0.0013046 | 0.94449 | 0.0013581 | 0.90122 | 0.0017906
100000 | 0.94243 | 0.0005666 | 0.94804 | 0.00058916 | 0.90639 | 0.00077757 | 0.03007% | 0.03211% | 0.03741% |  |  |  |  | 0.94178 | 0.00080495 | 0.94735 | 0.00083811 | 0.90536 | 0.001106
[7 empty] | -100% | -100% | -100% |  |  |  |  | 0.94243 | 0.0005666 | 0.94804 | 0.00058916 | 0.90639 | 0.00077757
[7 empty] | -100% | -100% | -100%
Relative uncertainty due to MC calculation
N | 100000
u_r | 0.1044%
recommanded value |  | 10000
the calculation uncertainty provided by the code in consistent with the observation

## Ni-63
The Stochastic implementation of TDCRPy is tested with its analystical implementation in the cases of pure beta emitting radionuclides
Convergence of the results with the number of points considered in the quenching function (parameter nE)
Analytical model, L=1 keV-1, kB=1e-5 cm/keV | [5 empty] | Ni-63 | Q | 66.98 | keV
nE | eff_D | Err
3 | 90.64% | 8.917%
4 | 89% | 6.955%
5 | 87.95% | 5.692%
6 | 87.22% | 4.808%
7 | 86.67% | 4.155%
10 | 85.65% | 2.929%
20 | 84.4% | 1.418%
50 | 83.63% | 0.4989%
80 | 83.46% | 0.2909%
100 | 83.41% | 0.2274%
200 | 83.31% | 0.1101%
500 | 83.25% | 0.04271%
1000 | 83.23% | 0.01973%
2000 | 83.22% | 0.008571%
5000 | 83.22% | 0.002205%
7000 | 83.22% | 0.0009248%
10000 | 83.22% | 0%
20000 |  | -100%
50000 |  | -100%
70000 |  | -100%
reference | 83.22%
recommanded value |  | 2000
Convergence of the result of the stochastic model with the results o the analytical model (N = number of Monte-Carlo trials) | [9 empty] | Analytical model results
Analytical model, L=1 keV-1, kB=1e-5 cm/keV | [9 empty] | eff | TDCRPy | TDCR17 restults | deviation
N | eff_S | u(eff_S) | eff_D | u(eff_D) | eff_T | u(eff_T) | dev S | dev D | dev T | eff_S | 0.81694 |  | #DIV/0!
10 | 76.05% | 9.644% | 77.04% | 10.51% | 62.18% | 12.73% | -6.912% | -7.426% | -11.11% | eff_D | 0.83224 | 0.8316 | 0.07674%
20 | 73.24% | 7.09% | 74.18% | 7.553% | 58.46% | 9.283% | -10.35% | -10.87% | -16.42% | eff_T | 0.69947 | 0.6989 | 0.08122% | 0.76048 | 0.096439 | 0.77044 | 0.10505 | 0.62178 | 0.12734
50 | 83.74% | 3.7% | 85.5% | 4.016% | 72.71% | 4.944% | 2.507% | 2.737% | 3.944% | TDCR | 0.84047 | 0.84043 | 0.00448% | 0.73243 | 0.070898 | 0.74176 | 0.075529 | 0.58459 | 0.092834
100 | 83.93% | 2.582% | 85.89% | 2.813% | 72.71% | 3.425% | 2.735% | 3.205% | 3.957% |  |  |  |  | 0.83742 | 0.036996 | 0.85502 | 0.040165 | 0.72706 | 0.049438
200 | 82% | 1.892% | 83.5% | 2.055% | 69.85% | 2.611% | 0.3725% | 0.3265% | -0.1371% |  |  |  |  | 0.83928 | 0.025823 | 0.85891 | 0.028133 | 0.72715 | 0.034248
500 | 81.54% | 1.259% | 82.99% | 1.378% | 70.11% | 1.647% | -0.1945% | -0.2848% | 0.2275% |  |  |  |  | 0.81998 | 0.018923 | 0.83496 | 0.020552 | 0.69851 | 0.026114
1000 | 81.87% | 0.8722% | 83.42% | 0.9517% | 70.25% | 1.156% | 0.2174% | 0.2315% | 0.4305% |  |  |  |  | 0.81535 | 0.012586 | 0.82987 | 0.013783 | 0.70106 | 0.016465
2000 | 82.19% | 0.6034% | 84.04% | 0.6547% | 70.23% | 0.8031% | 0.6062% | 0.9782% | 0.4061% |  |  |  |  | 0.81872 | 0.0087223 | 0.83417 | 0.0095173 | 0.70248 | 0.011556
5000 | 82.33% | 0.3827% | 83.92% | 0.4145% | 70.69% | 0.515% | 0.7734% | 0.8348% | 1.057% |  |  |  |  | 0.82189 | 0.0060343 | 0.84038 | 0.0065468 | 0.70231 | 0.0080312
10000 | 81.73% | 0.2779% | 83.22% | 0.3019% | 70.17% | 0.3683% | 0.04363% | -0.002859% | 0.3132% |  |  |  |  | 0.82326 | 0.0038267 | 0.83919 | 0.0041447 | 0.70686 | 0.0051504
20000 | 81.54% | 0.1974% | 83.03% | 0.2147% | 69.91% | 0.2611% | -0.185% | -0.2336% | -0.05482% |  |  |  |  | 0.8173 | 0.0027786 | 0.83221 | 0.0030194 | 0.70166 | 0.003683
50000 | 81.76% | 0.1231% | 83.3% | 0.1338% | 70% | 0.1645% | 0.08293% | 0.09064% | 0.07377% |  |  |  |  | 0.81543 | 0.0019735 | 0.83029 | 0.0021468 | 0.69908 | 0.0026113
100000 | 81.74% | 0.08752% | 83.27% | 0.09502% | 70.08% | 0.1164% | 0.05853% | 0.05277% | 0.1916% |  |  |  |  | 0.81762 | 0.0012314 | 0.83299 | 0.0013377 | 0.69998 | 0.0016449
[7 empty] | -100% | -100% | -100% |  |  |  |  | 0.81742 | 0.00087517 | 0.83268 | 0.00095023 | 0.70081 | 0.0011637
[7 empty] | -100% | -100% | -100%
Relative uncertainty due to MC calculation
N | 100000
u_r | 0.09961%
recommanded value |  | 10000
the calculation uncertainty provided by the code in consistent with the observation

## Ca-45
The Stochastic implementation of TDCRPy is tested with its analystical implementation in the cases of pure beta emitting radionuclides
Convergence of the results with the number of points considered in the quenching function (parameter nE)
Analytical model, L=1 keV-1, kB=1e-5 cm/keV | [5 empty] | Ca-45 | Q | 258 | keV
nE | eff_D | Err
3 |  | -100%
4 |  | -100%
5 |  | -100%
6 |  | -100%
7 |  | -100%
10 |  | -100%
20 | 96.46% | 0.2694%
50 | 96.29% | 0.09523%
80 | 96.25% | 0.05554%
100 | 96.24% | 0.0433%
200 | 96.22% | 0.02094%
500 | 96.21% | 0.008154%
1000 | 96.2% | 0.003772%
2000 | 96.2% | 0.00161%
5000 | 96.2% | 0.0004331%
7000 | 96.2% | 0.0001847%
10000 | 96.2% | 0%
20000 |  | -100%
50000 |  | -100%
70000 |  | -100%
reference | 96.2%
recommanded value |  | 500
Convergence of the result of the stochastic model with the results o the analytical model (N = number of Monte-Carlo trials) | [9 empty] | Analytical model results
Analytical model, L=1 keV-1, kB=1e-5 cm/keV | [9 empty] | eff | TDCRPy | TDCR17 restults | deviation
N | eff_S | u(eff_S) | eff_D | u(eff_D) | eff_T | u(eff_T) | dev S | dev D | dev T | eff_S | 0.95786
10 | 0.90153 | 0.055642 | 0.91413 | 0.053117 | 0.80851 | 0.10314 | -5.881% | -4.983% | -13.18% | eff_D | 0.96207 | 0.9667 | 0.4809%
20 | 0.95373 | 0.04428 | 0.95112 | 0.047603 | 0.94765 | 0.048625 | -0.4318% | -1.138% | 1.765% | eff_T | 0.93121 | 0.9357 | 0.4819% | 0.90153 | 0.055642 | 0.91413 | 0.053117 | 0.80851 | 0.10314
50 | 0.89719 | 0.036285 | 0.89383 | 0.039244 | 0.85925 | 0.045934 | -6.334% | -7.094% | -7.728% | TDCR | 0.96792 | 0.96793 | -0.001% | 0.95373 | 0.04428 | 0.95112 | 0.047603 | 0.94765 | 0.048625
100 | 0.98023 | 0.0092313 | 0.98506 | 0.009601 | 0.96153 | 0.01433 | 2.335% | 2.389% | 3.256% |  |  |  |  | 0.89719 | 0.036285 | 0.89383 | 0.039244 | 0.85925 | 0.045934
200 | 0.9686 | 0.0095345 | 0.97427 | 0.0095255 | 0.94742 | 0.012608 | 1.121% | 1.268% | 1.741% |  |  |  |  | 0.98023 | 0.0092313 | 0.98506 | 0.009601 | 0.96153 | 0.01433
500 | 0.95482 | 0.0077777 | 0.95682 | 0.0081875 | 0.93448 | 0.0097981 | -0.3177% | -0.5462% | 0.3507% |  |  |  |  | 0.9686 | 0.0095345 | 0.97427 | 0.0095255 | 0.94742 | 0.012608
1000 | 0.9683 | 0.0038737 | 0.97356 | 0.0039823 | 0.94213 | 0.005914 | 1.09% | 1.194% | 1.172% |  |  |  |  | 0.95482 | 0.0077777 | 0.95682 | 0.0081875 | 0.93448 | 0.0097981
2000 | 0.95493 | 0.0035886 | 0.95942 | 0.0036914 | 0.92677 | 0.0049487 | -0.3056% | -0.2758% | -0.4766% |  |  |  |  | 0.9683 | 0.0038737 | 0.97356 | 0.0039823 | 0.94213 | 0.005914
5000 | 0.95792 | 0.0021769 | 0.96208 | 0.0022546 | 0.93091 | 0.003042 | 0.005719% | 0.0004304% | -0.03256% |  |  |  |  | 0.95493 | 0.0035886 | 0.95942 | 0.0036914 | 0.92677 | 0.0049487
10000 | 0.95806 | 0.001563 | 0.96206 | 0.0016177 | 0.93245 | 0.0021363 | 0.02116% | -0.001665% | 0.1326% |  |  |  |  | 0.95792 | 0.0021769 | 0.96208 | 0.0022546 | 0.93091 | 0.003042
20000 | 0.95821 | 0.0010873 | 0.96269 | 0.001123 | 0.93139 | 0.0015063 | 0.03685% | 0.06414% | 0.01866% |  |  |  |  | 0.95806 | 0.001563 | 0.96206 | 0.0016177 | 0.93245 | 0.0021363
50000 | 0.95795 | 0.00069277 | 0.96206 | 0.00071841 | 0.93148 | 0.00095816 | 0.009609% | -0.001465% | 0.02855% |  |  |  |  | 0.95821 | 0.0010873 | 0.96269 | 0.001123 | 0.93139 | 0.0015063
100000 | 0.95644 | 0.0005004 | 0.96061 | 0.00051933 | 0.92942 | 0.0006877 | -0.1484% | -0.1516% | -0.1925% |  |  |  |  | 0.95795 | 0.00069277 | 0.96206 | 0.00071841 | 0.93148 | 0.00095816
[7 empty] | -100% | -100% | -100% |  |  |  |  | 0.95644 | 0.0005004 | 0.96061 | 0.00051933 | 0.92942 | 0.0006877
[7 empty] | -100% | -100% | -100%
Relative uncertainty due to MC calculation
N | 100000
u_r | 0.1044%
recommanded value |  | 10000
the calculation uncertainty provided by the code in consistent with the observation

## Tc-99
The Stochastic implementation of TDCRPy is tested with its analystical implementation in the cases of pure beta emitting radionuclides
Convergence of the results with the number of points considered in the quenching function (parameter nE)
Analytical model, L=1 keV-1, kB=1e-5 cm/keV | [5 empty] | Tc-99 | Q | 293.8 | keV
nE | eff_D | Err
3 |  | -100%
4 |  | -100%
5 |  | -100%
6 |  | -100%
7 |  | -100%
10 | 97.81% | 0.3781%
20 | 97.62% | 0.1844%
50 | 97.5% | 0.06529%
80 | 97.47% | 0.03802%
100 | 97.47% | 0.02971%
200 | 97.45% | 0.01443%
500 | 97.44% | 0.005562%
1000 | 97.44% | 0.002609%
2000 | 97.44% | 0.00115%
5000 | 97.44% | 0.0002899%
7000 | 97.44% | 0.0001201%
10000 | 97.44% | 0%
20000 |  | -100%
50000 |  | -100%
70000 |  | -100%
reference | 97.44%
recommanded value |  | 500
Convergence of the result of the stochastic model with the results o the analytical model (N = number of Monte-Carlo trials) | [9 empty] | Analytical model results
Analytical model, L=1 keV-1, kB=1e-5 cm/keV | [9 empty] | eff | TDCRPy | TDCR17 restults | deviation
N | eff_S | u(eff_S) | eff_D | u(eff_D) | eff_T | u(eff_T) | dev S | dev D | dev T | eff_S | 0.97138 |  | #DIV/0!
10 | 89.96% | 8.451% | 89.98% | 9.173% | 87.11% | 9.567% | -7.389% | -7.661% | -8.57% | eff_D | 0.97443 | 0.9724 | 0.209%
20 | 96.66% | 3.144% | 96.43% | 3.473% | 94.9% | 4.648% | -0.4913% | -1.039% | -0.4006% | eff_T | 0.9528 | 0.9464 | 0.6764%
50 | 97.44% | 1.617% | 97.81% | 1.748% | 95.47% | 2.345% | 0.3078% | 0.3725% | 0.1957% | TDCR | 0.9778 | 0.97326 | 0.4665%
100 | 96.81% | 1.408% | 97.34% | 1.412% | 94.88% | 1.792% | -0.3386% | -0.1067% | -0.4185% |  |  |  |  | 0.8996 | 0.084506 | 0.89978 | 0.091731 | 0.87115 | 0.095671
200 | 97.57% | 0.8129% | 97.83% | 0.8311% | 95.88% | 1.191% | 0.4497% | 0.3934% | 0.6308% |  |  |  |  | 0.96661 | 0.031438 | 0.96431 | 0.034733 | 0.94898 | 0.046481
500 | 96.64% | 0.6128% | 97.04% | 0.6185% | 94.38% | 0.8713% | -0.5139% | -0.4186% | -0.9403% |  |  |  |  | 0.97437 | 0.016169 | 0.97806 | 0.017479 | 0.95467 | 0.023449
1000 | 97.23% | 0.3904% | 97.59% | 0.3994% | 95.29% | 0.5562% | 0.09322% | 0.1514% | 0.01095% |  |  |  |  | 0.96809 | 0.014075 | 0.97339 | 0.01412 | 0.94881 | 0.017916
2000 | 97.05% | 0.2939% | 97.32% | 0.3058% | 95.22% | 0.4068% | -0.08537% | -0.1219% | -0.05867% |  |  |  |  | 0.97575 | 0.0081292 | 0.97826 | 0.0083106 | 0.95881 | 0.011914
5000 | 97.06% | 0.1831% | 97.38% | 0.188% | 95.14% | 0.2573% | -0.07814% | -0.0623% | -0.1493% |  |  |  |  | 0.96639 | 0.0061282 | 0.97035 | 0.0061852 | 0.94384 | 0.008713
10000 | 97.16% | 0.128% | 97.46% | 0.1324% | 95.33% | 0.1786% | 0.02299% | 0.01251% | 0.05396% |  |  |  |  | 0.97228 | 0.0039036 | 0.97591 | 0.0039937 | 0.95291 | 0.0055621
20000 | 96.96% | 0.09294% | 97.3% | 0.09579% | 94.97% | 0.1302% | -0.1793% | -0.1444% | -0.3264% |  |  |  |  | 0.97055 | 0.0029393 | 0.97324 | 0.0030584 | 0.95224 | 0.0040677
50000 | 96.97% | 0.05948% | 97.27% | 0.06143% | 95.06% | 0.08241% | -0.175% | -0.177% | -0.2323% |  |  |  |  | 0.97062 | 0.0018308 | 0.97382 | 0.0018801 | 0.95138 | 0.0025727
100000 | 97.04% | 0.04118% | 97.35% | 0.04249% | 95.11% | 0.05768% | -0.1049% | -0.09302% | -0.1817% |  |  |  |  | 0.9716 | 0.0012801 | 0.97455 | 0.0013237 | 0.95332 | 0.0017856
[7 empty] | -100% | -100% | -100% |  |  |  |  | 0.96964 | 0.00092938 | 0.97302 | 0.00095791 | 0.94969 | 0.0013018
[7 empty] | -100% | -100% | -100% |  |  |  |  | 0.96968 | 0.00059475 | 0.97271 | 0.0006143 | 0.95059 | 0.00082409
[14 empty] | 0.97036 | 0.00041182 | 0.97353 | 0.00042495 | 0.95107 | 0.00057677
Relative uncertainty due to MC calculation
N | 100000
u_r | 0.1044%
recommanded value |  | 10000
the calculation uncertainty provided by the code in consistent with the observation

## Sr-90
The Stochastic implementation of TDCRPy is tested with its analystical implementation in the cases of pure beta emitting radionuclides
Convergence of the results with the number of points considered in the quenching function (parameter nE)
Analytical model, L=1 keV-1, kB=1e-5 cm/keV | [5 empty] | Sr-90 | Q | 545.9 | keV
nE | eff_D | Err
3 | 99.21% | 0.5779%
4 | 99.08% | 0.454%
5 | 99.01% | 0.3738%
6 | 98.95% | 0.3165%
7 | 98.91% | 0.2745%
10 | 98.83% | 0.1943%
20 | 98.73% | 0.09478%
50 | 98.67% | 0.03358%
80 | 98.66% | 0.01955%
100 | 98.65% | 0.01526%
200 | 98.64% | 0.007364%
500 | 98.64% | 0.002707%
1000 | 98.64% | 0.001175%
2000 | 98.64% | 0.0004564%
5000 | 98.64% | 7e-05%
7000 | 98.64% | 0%
10000 |  | -100%
20000 |  | -100%
50000 |  | -100%
70000 |  | -100%
reference | 98.64%
recommanded value |  | 200
Convergence of the result of the stochastic model with the results o the analytical model (N = number of Monte-Carlo trials) | [9 empty] | Analytical model results
Analytical model, L=1 keV-1, kB=1e-5 cm/keV | [9 empty] | eff | TDCRPy | TDCR17 restults | deviation
N | eff_S | u(eff_S) | eff_D | u(eff_D) | eff_T | u(eff_T) | dev S | dev D | dev T | eff_S | 0.98487 |  | #DIV/0!
10 | 1 | 3.4191e-07 | 1 | 3.6968e-12 | 1 | 1.0257e-06 | 1.537% | 1.375% | 2.539% | eff_D | 0.98644 | 0.987 | -0.05676%
20 | 1 | 2.9332e-06 | 1 | 5.2978e-10 | 0.99999 | 8.7989e-06 | 1.536% | 1.375% | 2.538% | eff_T | 0.97524 | 0.9737 | 0.1581% | 1 | 3.4191e-07 | 1 | 3.6968e-12 | 1 | 1.0257e-06
50 | 0.9929 | 0.0042034 | 0.99748 | 0.0017774 | 0.98137 | 0.010795 | 0.8159% | 1.119% | 0.6285% | TDCR | 0.98865 | 0.98652 | 0.215% | 1 | 2.9332e-06 | 1 | 5.2978e-10 | 0.99999 | 8.7989e-06
100 | 0.98456 | 0.010055 | 0.98507 | 0.010653 | 0.97668 | 0.013495 | -0.0307% | -0.1392% | 0.1479% |  |  |  |  | 0.9929 | 0.0042034 | 0.99748 | 0.0017774 | 0.98137 | 0.010795
200 | 0.98598 | 0.0072246 | 0.9884 | 0.0071104 | 0.9797 | 0.0081033 | 0.1131% | 0.1985% | 0.4572% |  |  |  |  | 0.98456 | 0.010055 | 0.98507 | 0.010653 | 0.97668 | 0.013495
500 | 0.97953 | 0.0053246 | 0.98132 | 0.0054809 | 0.96981 | 0.0065991 | -0.5424% | -0.5187% | -0.5571% |  |  |  |  | 0.98598 | 0.0072246 | 0.9884 | 0.0071104 | 0.9797 | 0.0081033
1000 | 0.98474 | 0.0029257 | 0.9868 | 0.0029167 | 0.97401 | 0.0042347 | -0.0128% | 0.03611% | -0.1257% |  |  |  |  | 0.97953 | 0.0053246 | 0.98132 | 0.0054809 | 0.96981 | 0.0065991
2000 | 0.98715 | 0.0019928 | 0.98902 | 0.0019778 | 0.97909 | 0.0026767 | 0.2313% | 0.2619% | 0.3944% |  |  |  |  | 0.98474 | 0.0029257 | 0.9868 | 0.0029167 | 0.97401 | 0.0042347
5000 | 0.98239 | 0.001453 | 0.98412 | 0.0014968 | 0.97141 | 0.0020175 | -0.2515% | -0.2354% | -0.3924% |  |  |  |  | 0.98715 | 0.0019928 | 0.98902 | 0.0019778 | 0.97909 | 0.0026767
10000 | 0.98245 | 0.0010424 | 0.98412 | 0.0010765 | 0.97207 | 0.0014112 | -0.2449% | -0.235% | -0.3255% |  |  |  |  | 0.98239 | 0.001453 | 0.98412 | 0.0014968 | 0.97141 | 0.0020175
20000 | 0.98302 | 0.00072044 | 0.98459 | 0.0007409 | 0.97277 | 0.0009903 | -0.1875% | -0.1879% | -0.2529% |  |  |  |  | 0.98245 | 0.0010424 | 0.98412 | 0.0010765 | 0.97207 | 0.0014112
50000 | 0.98381 | 0.00044164 | 0.98544 | 0.00045283 | 0.97373 | 0.00061187 | -0.1075% | -0.1012% | -0.1552% |  |  |  |  | 0.98302 | 0.00072044 | 0.98459 | 0.0007409 | 0.97277 | 0.0009903
100000 | 0.98436 | 0.00030504 | 0.98602 | 0.00031323 | 0.9744 | 0.00042494 | -0.05123% | -0.04278% | -0.08573% |  |  |  |  | 0.98381 | 0.00044164 | 0.98544 | 0.00045283 | 0.97373 | 0.00061187
[7 empty] | -100% | -100% | -100% |  |  |  |  | 0.98436 | 0.00030504 | 0.98602 | 0.00031323 | 0.9744 | 0.00042494
[7 empty] | -100% | -100% | -100%
Relative uncertainty due to MC calculation
N | 100000
u_r | 0.1044%
recommanded value |  | 10000
the calculation uncertainty provided by the code in consistent with the observation

## Sr-89
The Stochastic implementation of TDCRPy is tested with its analystical implementation in the cases of pure beta emitting radionuclides
Convergence of the results with the number of points considered in the quenching function (parameter nE)
Analytical model, L=1 keV-1, kB=1e-5 cm/keV | [5 empty] | Sr-89 | Q | 1495.1 | keV
nE | eff_D | Err
3 |  | -100%
4 |  | -100%
5 |  | -100%
6 |  | -100%
7 |  | -100%
10 |  | -100%
20 | 99.58% | 0.02319%
50 | 99.56% | 0.007916%
80 | 99.56% | 0.004416%
100 | 99.56% | 0.003453%
200 | 99.55% | 0.00152%
500 | 99.55% | 0.0003729%
1000 | 99.55% | 0%
2000 |  | -100%
5000 |  | -100%
7000 |  | -100%
10000 |  | -100%
20000 |  | -100%
50000 |  | -100%
70000 |  | -100%
reference | 99.55%
recommanded value |  | 50
Convergence of the result of the stochastic model with the results o the analytical model (N = number of Monte-Carlo trials) | [9 empty] | Analytical model results
Analytical model, L=1 keV-1, kB=1e-5 cm/keV | [9 empty] | eff | TDCRPy | TDCR17 restults | deviation
N | eff_S | u(eff_S) | eff_D | u(eff_D) | eff_T | u(eff_T) | dev S | dev D | dev T | eff_S | 0.9951 |  | #DIV/0!
10 | 100% | 1.711e-08% | 100% | 0% | 100% | 5.132e-08% | 0.4928% | 0.4419% | 0.7205% | eff_D | 0.9956 | 0.9982 | -0.2605%
20 | 100% | 3.008e-13% | 100% | 0% | 100% | 9.025e-13% | 0.4928% | 0.4419% | 0.7205% | eff_T | 0.99285 | 0.9958 | -0.2966%
50 | 97.01% | 2.183% | 97.04% | 2.178% | 96.21% | 2.589% | -2.515% | -2.529% | -3.1% | TDCR | 0.99723 | 0.9976 | -0.03627% | 1 | 1.7106e-10 | 1 | 0 | 1 | 5.1318e-10
100 | 98% | 1.4% | 98% | 1.4% | 98% | 1.4% | -1.518% | -1.567% | -1.296% |  |  |  |  | 1 | 3.0084e-15 | 1 | 0 | 1 | 9.0248e-15
200 | 99.96% | 0.03161% | 99.99% | 0.005235% | 99.88% | 0.08964% | 0.4495% | 0.4361% | 0.5966% |  |  |  |  | 0.97006 | 0.021829 | 0.97042 | 0.021775 | 0.96207 | 0.025889
500 | 99.5% | 0.2376% | 99.61% | 0.2269% | 99.12% | 0.3333% | -0.009897% | 0.04866% | -0.1634% |  |  |  |  | 0.97999 | 0.014 | 0.98 | 0.014 | 0.97997 | 0.014
1000 | 99.39% | 0.202% | 99.46% | 0.1981% | 99.06% | 0.2644% | -0.1193% | -0.09763% | -0.2276% |  |  |  |  | 0.99957 | 0.00031614 | 0.99994 | 5.2349e-05 | 0.99877 | 0.0008964
2000 | 99.26% | 0.1681% | 99.34% | 0.1655% | 98.97% | 0.1975% | -0.2503% | -0.2177% | -0.3211% |  |  |  |  | 0.995 | 0.0023762 | 0.99608 | 0.0022686 | 0.99122 | 0.0033331
5000 | 99.41% | 0.08836% | 99.48% | 0.08655% | 99.08% | 0.1171% | -0.09921% | -0.08225% | -0.2016% |  |  |  |  | 0.99391 | 0.0020196 | 0.99463 | 0.0019808 | 0.99059 | 0.0026438
10000 | 99.39% | 0.06694% | 99.45% | 0.06613% | 99.12% | 0.08246% | -0.1214% | -0.1142% | -0.1645% |  |  |  |  | 0.99261 | 0.0016809 | 0.99343 | 0.001655 | 0.98966 | 0.0019749
20000 | 99.32% | 0.04994% | 99.39% | 0.0494% | 99.02% | 0.06124% | -0.1913% | -0.1746% | -0.2664% |  |  |  |  | 0.99411 | 0.0008836 | 0.99478 | 0.00086554 | 0.99084 | 0.0011712
50000 | 99.39% | 0.02937% | 99.46% | 0.02894% | 99.1% | 0.03663% | -0.1158% | -0.09886% | -0.1812% |  |  |  |  | 0.99389 | 0.00066944 | 0.99446 | 0.00066135 | 0.99121 | 0.00082464
100000 | 99.31% | 0.02224% | 99.39% | 0.02194% | 99% | 0.02755% | -0.1971% | -0.1751% | -0.2901% |  |  |  |  | 0.99319 | 0.00049943 | 0.99386 | 0.00049402 | 0.9902 | 0.00061239
[7 empty] | -100% | -100% | -100% |  |  |  |  | 0.99394 | 0.00029367 | 0.99462 | 0.0002894 | 0.99105 | 0.00036632
[7 empty] | -100% | -100% | -100% |  |  |  |  | 0.99314 | 0.00022237 | 0.99386 | 0.00021944 | 0.98997 | 0.00027551
Relative uncertainty due to MC calculation
N | 100000
u_r | 0.1044%
recommanded value |  | 10000
the calculation uncertainty provided by the code in consistent with the observation